Biological Incubator

SPX-70/100/150/250

Manual

Manufacturer : Kenton apparatus co.,ltd

Factory Add: NO.,88, Shibashe industrial zone, Xiamao, Shijing, Baiyun District,Guangzhou, Guangdong, China

Tel: 0086-20-36246586  36247961

Fax: 0086-20-61152082

For detail information please visit the our website: http://www.kentonchina.com/

Thanks for purchasing Kenton products, please read this manual carefully before using, finish reading, put it in convenient place for later use

- Application

SPX-70/100/150/250 series biochemical incubator is thermostatic equipment with precise and highly sophisticated cooling function. It is widely used in biological genetic engineering, medical treatment, health and epidemic prevention, medical test, agriculture and animal husbandry, aquaculture. And other scientific research institutions for plant cultivation, breeding test, bacteria, mould, microbial cultivation & preservation, BOD determination, etc.

2. Product Structure

3. Features

The incubator is made of high quality steel sheet with sharp appearance. Equip with large observation glass window and built in lighting lamp. Samples inside can be observed clearly. The inner chamber is made of stainless steel, equipped with adjustable shelves. Easy to clean and durable.

4. Working principle

The actual temperature is detected from temperature sensor which transferred into signal. The signal is transmitted through microcomputer and then sent to the heater and compressor to function towards required temperature.

- Technical Parameter

1.Volume:70L,100L,150L,250L；

2.Temp. range:5～50℃；

3.Temp. fluctuation:±1℃（70L,100L,150L）   ±1.5℃（250L）；

4.Temp. uniformity:±1.5℃（70L,100L,150L）   ±2℃（250L）；

5.Power voltage:AC 220V/50Hz；

6.Input power:385W（70L,100L）     420W（150L,250L）；

7.Working ambient temp.10~30℃   Relative humidity below 70%

8.Refrigeration:R134a；

9.Equipment classify: class I  type B

6. Installation

1.The equipment should be installed in ventilated place and avoid direct sunlight. The distance between equipment and wall have to be much than 10 CM

2.Keep the incubator balance via adjustable screw. (Refer to technical diagram)

3.AC 220V/50HZ power supply, in order to be safe, please connect protective conductor thermal

4．Please install this equipment in the place where is far away from fire source

7.Control Panel

Indicator and button instruction

1．"heat"indicator:turns on if there is heating output , or turns off

2．"cool" indicator:turns on if there is cooling output , or turns off

3．"alarm" indicator:turns on if temperature is abnormal or over-temperature alarm, or turns off

4．"defrost" indicator:  no defrost function under intermittent mode

5．"lighting" indicator:turns on when lighting lamp is on, or turns off

6．"set temp." indicator:turns on if the right screen displays set temperature, or turns off

7．"set time" indicator:turns on if right screen displays running time, or turns off

8．"set"button:set or view temperature, time and other parameters

9．"◄"shift button:press this button to modify set value in setting mode

10．"▼"decrease / rerun:press this button to decrease set value in setting mode, hold this button decrease continuously. After running is over, in non-setting mode, hold this button to reboot.

11．"▲"increase:increase set value in setting value, hold this button increase continuously

12．"lighting"button:lighting lamp button, press this button to turn on or turn off the lighting lamp

8．Operation

1．After power is connected, left screen displays "index number and version number", right screen displays "the max setting temperature" . It will access home interface after 3 seconds.

2．View and set temperature and time of constant temperature

1）if without timer

Press "set" button to access temperature setting mode, left screen displays "SP" right screen displays setting temperature(it will flash from single digit), modify required setting value via shift ,increase or decrease button; re-press "set" button to log out this mode, the modified data will be saved automatically.

2）if with timer

Press "set" button to access temperature setting mode, left screen displays "SP" right screen displays setting temperature (it will flash from single digit), modify required setting value via shift ,increase or decrease button; re-press "set" to access timing setting mode , left screen displays "ST", right screen displays timing time (it will flash from single digit);re-press "set" button to log out this setting mode, the modified data will be saved automatically.

When timing is "0", it means no timer, equipment keeps running, right screen displays setting temperature; when timing time is not "0", right screen displays running time or setting temperature. If display running time, the decimal point of the single digit which is on the right screen is on, when measured temperature reaches setting temperature, the timer start timing, the decimal point of the single digit which is on the right screen is flashing, timing time is over and stop running, the right screen displays "End", beeper beeps, it will stop beeping after 1 minute. After running is over, hold "decrease" button for 3 seconds to reboot

Notice : if change setting temperature during timing, the equipment will start timing from 0

3. Sensor alarms abnormally

If the left screen displays "---", it means that there is something wrong with temperature sensor or equipment or temperature is in outside of temperature range, equipment will disconnect heating output automatically, beeper beeps, alarm indicator is on, please check the temperature sensor and other connections.

4. When oven-temperature alarm sets off, beeper beeps, "ALM" is on. if oven-temperature alarm because of setting temperature changed, "ALM" is off and without beep

5. Press any button to mute if beeper beeps

6. Without press any button within 1 minute in setting mode, it will be back to home interface automatically.

9. Internal Parameter Viewing And Setting

Hold set button for 3 seconds, the left screen displays "Lc", right screen displays password, you can modify required value via "increase" " decrease" or "shift" button . re-press "set" button , if the password is incorrect, equipment will be back to home interface, if the password is correct, it will access temperature setting mode. re-press "set" button modify parameters of. Re-hold "set" button for 3 seconds to log out this mode, the data will be saved automatically.

Internal Parameter -1

| Parameter code | Parameter | Parameter instruction | （range）factory default |
| --- | --- | --- | --- |
| Lc- | Password | When "Lc=3"view or set parameters | 0 |
| AL- | Over-temperature alarm | When measured temp. >setting temp. +AL", alarm indicator is on, beeper beeps , disconnect heating output When | (0～20.0℃) 5.0 |
| Ct- | Compressor delay started function | Compressor starts delay protection time, the mixed interval time from compressor stops to compressor reboots | (0～600 seconds) 60 |
| uP- | Compressor starts threshold | When compressor works in intermittent mode If measured temp. ≥ setting temp.+ uP" And delay time is up, start compressor Notice : this parameter is only available when start or stop compressor manually(Internal Parameter-2 , Crc=1) but it is unavailable in automatic mode | (-1.9～2.0℃) 0.4 |
| dn- | Compressor closes threshold | When compressor works in intermittent mode If measured temp. ≤setting temp. +dn" turn off compressor Notice : this parameter is only available when start or stop compressor manually(Internal Parameter-2 , Crc=1) but it is unavailable in automatic mode | (-2.0～(uP-0.1)) 0.2 |
| Lt- | Light lamp delay off function | When turn on light lamp, it will be off automatically after delay Lt .if "Lt=0"，delay is unavailable , you have to turn off light lamp manually | (0～30 mins) 2 |
| T- | Control cycle | Heating control cycle | (1～30seconds ) 5 |
| P- | Proportion | Timing proportion adjustment | (1～50.0) 16.0 |
| I- | Integral time | Integral acting in regulation | (1～2000 seconds ) 600 |
| d- | Differential time | Differential acting in regulation | (0～1000 seconds ) 150 |
| Pb- | Zero adjustment | Modify error margins of sensor (low temp.) during test Pb=actual temp.-test temp.) | (-12.0～12.0℃) 0 |
| PK- | Full scale adjustment | Modify error margins of sensor(high temp.) during test PK=1000*(actual temp.-test temp.)/ test temp. | (-999～999) 0 |

Internal Parameter -2

| Parameter code | Parameter | Parameter instruction | (range） factory default |
| --- | --- | --- | --- |
| Lc- | Password | When "Lc=9"view or set parameter | 0 |
| ndc- | Working way of compressor | 0:compressor only works in intermittent mode 1:compressor will work in balanced mode or in intermittent mode according to CP (details as below ) 2:compressor will work in balanced mode or in intermittent mode according to Htd (details as below ) | (0～2) 0 |
| CP- | The fixed switch point of compressor running mode | When ndc=1 If setting temperature ≥CP, compressor works in intermittent mode, on the contrary, it will work in balanced mode | (0～100.0℃) 30.0 |
| Htd | The auto-switch point of compressor running mode | When ndc=2 If setting temperature ≥ ambient temperature + Htd， compressor works in intermittent mode, on the contrary, it will work in balanced mode | (-50.0～50.0℃) 0.0 |
| Crc | Start/stop mode of compressor | When compressor works in intermittent mode 0: the compressor will be started or closed automatically (depend on ambient temp. and setting temp.) 1:start or close the compressor manually (depend on up and dn which on the Internal Parameter-1 | (0～1) 0 |
| Cnp | prohibited working temp. of compressor | When setting temperature ≥Cnp, compress stops working Notice : this running mode is in priority, it means the equipment will proceed with this first instead of balanced mode or intermittent mode | (0～100.0℃) 42.0 |
| ndt | Timer mode | 0:no timer； 1:with timer, when reach setting temperature（start timing），right screen displays running time 2:if without timer, right screen displays running time | (0～2) 0 |
| Hn- | Timer unit | 0:minutes ；1:hour | (0～1) 0 |
| nP- | The max power | The max power percentage of heating output | (0～100%) 100 |
| Co- | Disconnect heating output margins | When measured temperature ≥setting temperature +Co, disconnect heating output | (0～20.0℃) 5.0 |
| SPL- | The mix setting temp. | The mix setting temperature | (-15.0～5.0℃) 0.0 |
| SPH- | The max setting temp. | The max setting temperature | (5.0～100.0℃) 60.0 |

Internal Parameter -3

| Parameter code | Parameter | Parameter instruction | (range)factory default |
| --- | --- | --- | --- |
| Lc- | Password | When "Lc=18"view ambient temperature |  |
| Ht- | Ambient temp. | surrounding temperature of the equipment |  |

Notice: These Internal Parameters had been adjusted well , it is not necessary to re - adjust again

10. Maintenance

1.Keep upright when moving this biochemical incubator

2.In order to prolong the life of use of equipment, Do not change setting value frequently

3.This equipment is equipped with two fuses, if there is something wrong with this equipment when running, please disconnect power supply first, check the fuse whether is undamaged then check other parts

4.In order to maintain the appearance of equipment, Do not clean the equipment appearance with corrosive solution and clean inner chamber with dry cloth and alcohol

5.If set this equipment aside , please keep inner chamber dry and disconnect power supply

6.In order to make sure uniform temperature , frequently check fan whether works normally. The samples should not be overloaded in the inner chamber.

7.In order to make sure good temperature effect, Do not touch or collide the temperature sensor

8. This equipment is equipped with cooling function, which is cooling mode. When cooling stops, it is normal that if there is pull-in of magnet valve or refrigerant backflow sound

11. Trouble Shooting

| Troubleshooting | Handing |
| --- | --- |
| Screen displays (----) | ·There is something wrong with temperature sensor , check the temperature sensor , model : PT100 |
| Temperature can’t reach setting value | ·Setting temperature is lower than display temperature Check whether the heater is damaged Check the equipment board and other connections |
| Temperature rises slowly | Check whether the fan stops working fan is broken something wrong with relative connections |
| Screen displays nothing | check whether power switch is on check whether the plug cord is 220V check failures according to wiring layout |

11. Wiring Layout

12. Packing List

| NO. | Name | Quantity | Remarks |
| --- | --- | --- | --- |
| 1 | finish product | 1 |  |
| 2 | instruction manual | 1 |  |
| 3 | certification | 1 |  |
| 4 | quality warranty card | 1 |  |
| 3 | shelf | 2（ 70L） 2（100L） 2（150L） 2（250L） |  |

After-sales service

- One year warranty，lifetime service and this product is warranted for a period of twelve months from the date of original purchase, please retain your receipt to establish proof of purchase.

- Do not dismantle, repair or refit equipment without authorization and guidance from our company. This warranty does not extend to the repair or refit of any products.

- This warranty does not cover any damages which is caused by improper operation，improper power output，or improper cable connection.

- This warranty applies only to products sold in china. If the product is located outside in china, during the warranty period, we will offer the parts for free. but not bear freight charge. If outside of warranty period, you need to pay for the parts and the freight charge.

- The warranty does not cover consumable parts, such as light lamp or glass door and so on.